[image: image1.jpg]SANXIN




Model MP523 pH/ISE Meter

Operation Manual

Shanghai San-Xin Instrumentation, Inc.

MP500 Series of Electrochemical Meters

Ⅰ. Multi-parameter Meter

Model MP551pH/mV/ISE/Cond/DO Meter

Ⅱ. Double Parameter Meters

1. Model MP521 Lab pH/Cond Meter (pH: ±0.01 pH; Conductivity: ±1.0% FS)

2. Model MP522 precision pH/Cond Meter (pH: ±0.002 pH; Conductivity:±0.5% FS)

3. Model MP523 pH/ISE Meter

4. Model MP525 pH/DO Meter

5. Model MP526 Cond/DO Meter

Ⅲ. Single Parameter Meters

1. MP511 Lab pH Meter（±0.01 pH）

2. MP512 Precision pH Meter（±0.002 pH）

3. MP513 Lab Conductivity Meter（±1.0% FS）

4. MP515 Precision Conductivity Meter（±0.5% FS）

5. MP516 DO Meter

6. MP517 Sodium Ion Concentration Meter

7. MP518 Calcium Ion Concentration Meter

8. MP519 Fluoride Ion Concentration Meter

Table of Contents

[image: image2.png]



1. Brief Introduction ---------------------------------------------------------------------------------------

2. Technical Parameters -----------------------------------------------------------------------------------

3. Instructions to the Meter -------------------------------------------------------------------------------

3.1. LCD Display -----------------------------------------------------------------------------------------

3.2. Operation Keys --------------------------------------------------------------------------------------

3.3. The Storage, Recall and Elimination of the Measuring Information -------------------------

3.4. RS-232 Communication ----------------------------------------------------------------------------

3.5. Sockets ------------------------------------------------------------------------------------------------

4. pH Measurement ----------------------------------------------------------------------------------------

4.1. Preparation Work -------------------------------------------------------------------------------------

4.2. Electrode Calibration -------------------------------------------------------------------------------

4.3. Sample Test -------------------------------------------------------------------------------------------

4.4. Parameter Setting ------------------------------------------------------------------------------------

4.5. Considerations ----------------------------------------------------------------------------------------

4.6. The Self-diagnose Information --------------------------------------------------------------------
5.
mV and ORP Measurement --------------------------------------------------------------------------

5.1. Sample Test -------------------------------------------------------------------------------------------

5.2. Considerations ---------------------------------------------------------------------------------------

5.3. Parameter Setting ------------------------------------------------------------------------------------

6. Ion Concentration Measurement -----------------------------------------------------------------------

6.1. Preparation Work ------------------------------------------------------------------------------------

6.2. Parameter Setting ------------------------------------------------------------------------------------

6.3. Electrode Calibration --------------------------------------------------------------------------------

6.4. Sample Test -------------------------------------------------------------------------------------------

6.5. Considerations ---------------------------------------------------------------------------------------

7. Using Guide on Model 901 Intelligent Stirrer -------------------------------------------------------

7.1. Technical Specifications ----------------------------------------------------------------------------

7.2. Function Features ------------------------------------------------------------------------------------

7.3. Using Method ----------------------------------------------------------------------------------------

7.4. Considerations ---------------------------------------------------------------------------------------

8. Meter’s Complete Kit -----------------------------------------------------------------------------------

9. Warranty --------------------------------------------------------------------------------------------------

ChartⅠ Meter Parameter Setting Schedule

ChartⅡ Meter Restore to Producer Setting Schedule

Chart Ⅲ Code Icon and Abbreviation Schedule
1. Brief Instruction:

[image: image3.png]



Thanks for buying and using the model MP523 pH/ISE Meter (the following called “meter” in short).

Before using this meter, please read the operation manual carefully in order to help use and maintain it correctly. On the basis of improving instrument of performance constantly, we reserve the right of changing the content of this manual and accessories in case of not notifying in advance.

This meter is a perfect combination with the most advanced electronic technology, sensor technology and software design, it is the best lab precision pH, mV and ion concentration meter with the highest performance and the lowest cost. It is suitable for the trade such as the petroleum, chemical industry, medicine, power plant and environmental protection, etc., especially has more extensive application in the fields of education and scientific research.

Built-in microprocessor chip, beautiful appearance, multi-functional and easy to use, this meter has the following prominent features:

1.1 Meter meets with the requirement of international GLP standards. Its features up to automatic calibration, automatic temperature compensation, data storage, timing measurement, RS-232 output, clock display, function setting and automatic self-diagnose etc.

[image: image4.png]


1.2. Adopts digital filter and step slipping technology to intelligently improve meter’s response speed and result accuracy. “     ” will appear when reading to be stable.

1.3. Automatically recognize 13 kinds of pH standard buffer solution. User can choose anyone from four series of buffer solutions: Europe & U.S.A. series, NIST series, CHINA series, and Self-define solution.

1.4. High purified water and ammonia added purified water pH measuring mode can be set up, for this two special pH measuring modes, expect the general slope temperature compensation, also added function of solution temperature compensation to make measuring accuracy to be more improved, especially suitable for use in fields of electric power and petrochemical etc.

1.5. With EH redox potential measuring mode can directly display ORP value which against potential of standard hydrogen electrode.

1.6. There are 10 kinds of common ions can be selected in the ion concentration measuring mode, user also can self-define other ions.

1.7. Self-setting 2 kinds of calibration solution, automatic calibration, and can measure sample’s ion concentration directly.

1.8. Meter’s circuit board adopts SMT film-covering technology to improve meter’s production reliability.

1.9. Meter has the only product series number.

1.10. Dustproof and waterproof meter meets IP54 rating. All sockets are seal protected by the silica gel caps.

2. Technical Parameters:

[image: image5.png]



2.1. pH:

	Measuring range
	(-1.999 to 19.999) pH

	Resolution
	0.1/0.01/0.001 pH

	Accuracy
	Meter: ±0.002 pH, connect meter with electrode:±0.01pH

	Input current
	≤1×10-12 A

	Input impendance
	≥3×1012 Ω

	Stability
	±0.002 pH/3h

	Temp. compensation range
	(0 to 100) ℃（automatic or manual）


2.2. mV:

	Measuring range (mV/ORP/EH)
	-1999.9 mV to 0 to 1999.9mV

	Resolution
	0.1mV

	Accuracy
	±0.03% FS


2.3. Ion Concentration:

	Measuring range
	0.00 to 14.00pX

	Display unit
	pX, mol/L, mg/L, ppm

	Accuracy
	Meter:±0.5 % reading(monovalent), ±1.0 % reading (bivalent)

	Temp. compensation range
	0 to 60 ℃（automatic or manual）


2.4. Temperature:

	Measuring range
	-10℃ to 110℃

	Resolution
	0.1℃

	Accuracy
	5 to 60℃ range: ±0.4℃  other range: ±0.8℃


2.5. Other Technical Parameters:

	Data storage
	900 groups

	Storage content
	Measuring value series number, measuring value, temperature, ATC or MTC state, measuring date and time

	Power
	DC9V/300mA

	Communication connector
	RS-232

	Size and weight
	160 x 190 x 70mm/880g

	Quality and safety certification
	ISO9001:2000, CE and CMC


2.6. Working Condition:

	Environment temperature
	5 to 35 ℃（0.01 grade） 15 to 30 ℃（0.001 grade）

	Environmental humidity
	≤75%

	IP rating
	IP54 Dustproof and waterproof


3. Instructions to the Meter:

[image: image6.png]



3.1. LCD Display:

[image: image7.png]



① —— Parameters mode icon

② —— Measuring value

③ —— Time, date and the indication icons for special measuring state
④ —— Measuring unit
⑤ —— Temperature compensation state icon
ATC — automatic temperature compensation

MTC — manual temperature compensation

⑥ —— Serial number and icon as measuring value to be stored and recalled

M+ — measuring value to be stored icon

RM — reading to be recalled icon; The left digit means serial number
⑦ —— Temperature measuring value and unit
⑧ —— RS-232 communication icon. When this icon appears, means meter has been connected with computer
⑨ —— Timing measuring icon
⑩ —— Electrode calibration indicate icon

[image: image8.png]-



—— Measuring value to be stable icon

[image: image9.png]


—— Ion symbol

[image: image10.png]


3.2. Operation Keys:

[image: image11.png]Picture (4-4)



The meter has 8 operation keys in all.
3.2.1. <      > — Switch key
3.2.2. < CAL > — Calibration key, press the key make meter enters into calibration mode.

3.2.3. < MODE > — Function key

[image: image12.png]


[image: image13.png]


[image: image14.png]


(a) Short-time press（time<1.5s）switch the measuring parameters, the meter will display     →      →      (ion concentration) .
(b) Depress (time >2s) to enter into the parameter setting mode P1, and again short-time press, will in turn display P2, P3…
3.2.4. < UNIT > —Unit and parameter mode key

[image: image15.png]


(a) When in     mode, press the key to change pH resolution:0.001→0.01→0.1 pH;

[image: image16.png]


[image: image17.png]Picture (6-6)



[image: image18.png]


[image: image19.png]F-Y4



(b) When in     mode, press the key to in turn change the parameter mode     →→     ;
[image: image20.png]


(c) When in      mode, press the key to in turn change unit symbol: pX→mol/L→mg/L→ppm;

(d) Press key to choose modified position of parameter when in the parameter setting state.
3.2.5. < ENTER > — Entrance key, when in the state of calibration or parameter setting, press the key to confirm. Then the meter enters into measuring state.
3.2.6. < ▲ >and < ▼ > — Increase and decrease key.

When in MTC state, press the key to increase or decrease the temperature value. Short-time press to alter 0.1℃ each time, while long-time press, the number will alter rapidly. When in the state of parameter setting, press the key to alter the number or ON/OFF state.

3.2.7. < M+/RM > — The compound key of memory and recall, short-time press (press time less than 1.5s) to save the measuring data, long-time press (press time more than 2s) to recall the saved measuring data.
3.3. The Storage, Recall and Elimination of the Measuring Information:

3.3.1. Store the measuring information:

[image: image21.png]


(a) In the measuring mode, when the measuring data is stable and appear the “     ”, short-time press < M+/RM > key, LCD will display “ M+ ” icon and storage serial number, and meanwhile memory all the measuring information including time, date, serial number, measuring value, measuring unit, temperature, temperature compensation state and so on. Meter can separately store 300 groups of measuring information in the mode of pH, mV and ion concentration, can totally store 900 groups.
[image: image22.png]


(b) When setting the timing measuring function, LCD will display the “     ” icon, and meter will do timing measurement according to the time intervals set, and stores the measuring information;
3.3.2. Recall measuring information:

(a) Under the measuring mode, depress the < M+/RM > key, meter will recall the last stored information, and the storage number and “RM” icon will appear in the lower right corner of the LCD, and the complete measuring information, measuring time and date will be displayed alternatively on the upper right corner of LCD. Again press < ▼ > or < ▲ > key, meter will in turn recall all the measuring information, depress < ▼ > or < ▲ > key can quickly query measuring information under other serial number;
(b) In the recalling mode (there are “ RM ” and storage serial number in the lower right corner of the LCD), press < ENTER > key to return to the measuring mode.
3.3.3. Eliminate the stored measuring information:

[image: image23.png]


In the recalling mode, depress the < ENTER > key for 5 seconds, LCD will appear “      ” for 2 seconds. It means the storage has been eliminated, and returns to measuring mode.
3.4. RS-232 Communication:

[image: image24.png]


3.4.1. This meter adopts MP523 communication software which can achieve RS-232 communication function. The requirements for computer of this application software are: PC can stable operate the Windows XP operating system (Microsoft Excel 2000 or much higher version has been installed in PC). According to different selection of resolution setting (1280 x 1024 or 1280 x 800) to install MP523 communication software by following the indications. Open the MP523 communication software and connect meter, the LCD will display “    ” RS-232 communication icon. Meanwhile all the stored information including measuring value, measuring unit, temperature, temperature compensation state (ATC or MTC) and the last calibration information (calibration time, date and solution) will be transmitted to the computer.

[image: image25.png]


3.4.2. Again press < M+/RM > key after meter being connected with computer, or  timing measuring function has been set, all the measuring information will be transmitted to computer through RS-232 and won’t be stored in meter. And there will also display a curve graph of measurement value changing with time on the interface of computer.
3.4.3. User can do operation such as analyzing, counting or printing etc. for all measuring information in “Microsoft Excel” file through the key “Lead out” in the interface of computer.

[image: image26.png]


3.5. Sockets:
3.5.1. REF — Reference electrode socket

3.5.2. pH/mV/ISE —pH electrode, ion selective electrode and ORP electrode sockets（BNC socket）
3.5.3. TEMP — Temperature electrode socket（RCA microphone socket）

3.5.4. RS232 — RS-232 communication connector socket

3.5.5. DC9V — DC9V power socket, Ф2.5, inner “+ ”outside “ - ”.

4. pH Measurement:

[image: image27.png]U-EF




[image: image28.png]En




4.1. Preparation Work:
4.1.1. Switch in power, press <      > key to turn on.
[image: image29.png]


4.1.2. Short-press < MODE > key to switch to      mode.
4.1.3. Install electrode holder and debug.
4.1.4. Install the pH electrode into meter’s socket.
4.2. Electrode Calibration:
[image: image30.png]


4.2.1. Press < CAL > to enter into calibration mode, LCD display the twinkling “        ” indicate to enter into the first point calibration.
[image: image31.png]


[image: image32.png]


4.2.2. Wash the pH electrode in purified water and dry it, then immerge it into the pH7.00 buffer solution, rock the electrode holder and then still, waiting for the data stable and appear “     ”, then again press key< CAL >, the LCD will appear a twinkling 7.000 pH, calibration finishes after several seconds and appear a stable pH value and a twinkling “        ”, indicates the first point calibration has been finished and enters into the second point calibration.
[image: image33.png]U-EF



[image: image34.png]


4.2.3. Wash the pH electrode in purified water and dry it, then immerge it into the pH4.00 buffer solution, rock the electrode holder and then still, waiting for a stable data and appear “     ”, then again press key < CAL >, the LCD will appear a twinkling 4.005pH, calibration finishes after several seconds, LCD will appear a stable pH value and a twinkling “       ”, indicates the second point calibration has been finished and enters into the third point calibration.
[image: image35.png]En




[image: image36.png]ORP




[image: image37.png]mV




4.2.4. Wash the pH electrode in purified water and dry it, then immerge it into the pH10.01 buffer solution, rock the electrode holder and still, waiting “     ” for the data stable and appear, then again press key <CAL>, LCD will appear a twinkling 10.012 pH, calibration finishes after several seconds, appear a stable pH value and “              ” three calibration indication icon, see picture (4-1), indicates the three-point calibration has been finished and entered into measuring mode.

4.2.5. Note:

[image: image38.png]En




(a) This meter can adopt random one-point, two-point or three-point automatic calibration, after the first point calibration (see item 4.2.2.), press < ENTER > to confirm one-point calibration and enters into measuring mode. The indication icon “    ” for one-point calibration will appear on the lower left corner of LCD. When the measuring accuracy is ≤±0.1pH, user just need to choose one kind buffer solution to take one-point calibration according to the measuring range.
[image: image39.png]mV




[image: image40.png]mV




(b) After the second point calibration, (see item 4.2.3.), press < ENTER > to confirm two-point calibration and enter into measuring mode. The indication icon “          ” for tow-point calibration will appear on the lower left corner of LCD. User can choose pH4.00 and pH7.00 to calibrate if the measurement is just within the acidity range and choose pH7.00 and pH10.01 to calibrate if just within the alkalinity range.
(c) User should choose three-point calibration so as to reach a more accurate measurement if the measuring range is wider or if the electrode has been used for long or has ageing phenomenon. As to the new electrode which be used for the first time, it must be calibrated by three-point calibration to keep the unanimity of the meter slope adjustment with the pH electrode.
4.3. Sample Test:
[image: image41.png]ORP




Immerge pH electrode into the sample solution after washing and dry it, rock the electrode holder and still, when the LCD appears the icon “     ” to take the reading after displaying value to be stable.

Note:

(a) The closer the temperature of the sample solution and the calibration solution, the more accurate measuring value will be acquired. This is called the pH equal temperature measuring theory.

(b) Meter equips with 2pcs of pH electrodes, for generally using, suggest using 201T-M plastic pH/ATC three-in-one combination electrode. And suggest using 2503-C glass pH combination electrode for high accuracy measurement or measurement for solution which is weaker ionic strength, high viscosity and feculent and colloidal solution.

4.4. Parameter Setting:

4.4.1. pH measuring parameter setting schedule（chart（4-1））

Chart（4-1）

	Prompt Mark
	Parameter Setting Items
	Code
	Parameters

	P1
	pH buffer solution series selection
	[image: image42.png]



	USA(Europe & U.S.A series)

NIS (NIST series)

CH (China series)

CUS (Self-define)

	P2
	Purified water pH temperature compensation setting
	[image: image43.png]Frrc




	OFF-On (shut-set)

	P3
	Ammonia added purified water pH temperature compensation setting
	[image: image44.png]Frrc




	OFF-On (shut-set)

	P4
	Time for timing measuring setting
	
	0 to 99 minutes

	P5
	Temperature unit setting
	
	℃ ℉

	P6
	Date setting
	Date
	Month / Day / Year

	P7
	Time setting
	Time
	Hour / Minute

	P8
	Restore to producer setting
	
	OFF-On ( shut-set)


[image: image45.png]


4.4.2. pH buffer solution series selection（P1）

(a) Depress < MODE > key, meter enters into P1 mode: see picture（4-2）.

(b) Press < ▲ > or < ▼ > key to choose buffer solution series:

[image: image46.png]


(Europe & U.S.A series) — 1.679, 4.005, 7.000,10.012 and 12.454 pH

[image: image47.png]


(NIST series) — 1.679, 4.006, 6.865, 9.180 and 12.454 pH

[image: image48.png]


(China series) — 1.680, 4.003, 6.864, 9.182 and 12.460 pH

(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

(d) The producer setting is U.S.A (Europe & U.S.A series).

4.4.3. pH buffer solution “Self-define” setting:

For some special calibration solution, user can set in the “Self-define” mode, but must pay special attention that the “Self-define” parameters are the values under some fixed temperature, and meter must be calibrated under the same temperature, otherwise there might be a bigger error. The following are the steps of setting the “Self-define”.

[image: image49.png]Picture (5-1)



[image: image50.png]Picture (4-10)



(a) Depress < MODE > key to enter into P1, press < ▲ > or < ▼ > key to choose the solution series as the “     ”, then press < CAL > to enter into the setting interface, show as picture (4-3), the upper right corner “        ” means the setting of the first calibration point.
(b) Press < UNIT > key, the number will move to the right and flash one by one, press < ▲ > or < ▼ > when the number flashing to alter the number, and set number follow this operation. Picture (4-3) is the whole procedure for one-point, two-point and three-point calibration setting.

[image: image51.png]-1
(N[N]

£




[image: image52.png]


[image: image53.png]


Note: If choose “      ” setting, then “      ” icon will appear on the upper right corner of LCD (time display part).

4.4.4. Purified water pH temperature compensation setting（P2）

[image: image54.png]-2



(a) Short press < MODE > key in the mode P1, the meter enters into mode P2, see picture（4-4）.

[image: image55.png]neE



[image: image56.png]neE



(b) Press < ▲ > or < ▼ > to choose “    ”（purified water pH temperature compensation setting）or “     ”（shut）.

(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

(d) The producer setting is “ OFF ”.

[image: image57.png]mV




Note: There will appear“       ” icon in the upper right corner of the LCD if purified water temperature compensation function was set.

4.4.5. Ammonia added purified water pH temperature compensation setting（P3）
(a) Short press < MODE > in mode P2 to enter into mode P3, see picture（4-5）.

[image: image58.png]ISE




[image: image59.png]


[image: image60.png]neE



(b) Press < ▲ > or < ▼ > key to choose “    ”（ammonia added purified water pH temperature compensation setting）or “     ”（shut）.

(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

(d) The producer setting is “ OFF ”.
[image: image61.png]neE



Note: If set the ammonia purified water pH temperature compensation function, there will appear “      ” icon in the right upper corner of the LCD when in the measuring mode.

[image: image62.png]


4.4.6. Time for timing measuring setting（P4）

(a) Short press < MODE > key in mode P3 to enter into mode P4. See picture（4-6）.

[image: image63.png]Picture (4-9)




(b) Press < UNIT > key, the “   ” will move rightward in turn and flash. Press < ▲ > or < ▼ > to alter the number.
(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

(d) The producer setting is “ 0 ” second.

[image: image64.png]


[image: image65.png]


Note: in the form “         ”, the number in the front of “:” is minute, and the maximum setting is 99, in the return of “ :” is second, the maximum setting is 59. After setting the timing measurement mode, LCD will display the “    ” icon.

[image: image66.png]


4.4.7. Temperature unit ℃/℉ setting（P5）

(a) Short press < MODE > key in mode P4 to enter into mode P5, see picture（4-7）.

(b) Press < ▲ > or < ▼ > key to choose temperature unit: ℃or℉.

(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

4.4.8. Date setting（P6）

(a) Short press < MODE > key in mode P5 to enter into mode P6, see picture（4-8）.

[image: image67.png]


(b) Press < UNIT > key, the number will move rightward and flash, press < ▲ > or < ▼ > key to alter the number. The upper right is month-day and lower right is year.
(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

[image: image68.png]


4.4.9.Time setting（P7）

(a) Short press < MODE > key in mode P6 to enter into mode P7, see picture (4-9).

(b) Press < UNIT > key, the number will move rightward and flash, press < ▲ > or < ▼ > key to alter the number.
(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

[image: image69.png]


4.4.10. Restore to producer setting (P8)

(a) Short press < MODE > key in mode P7 to enter into mode P8, see picture (4-10).

[image: image70.png]


(b) Press < ▲ > key to choose “    ”, means that all the pH measuring parameters have been restored to the producer setting mode, and restore to measuring mode after 2 seconds.

Special Note: When the parameter setting is compatible with the producer setting, the pH measuring interface will show as picture（4-1）,time is displayed on the upper right corner of the LCD. When the parameter setting is not compatible with the producer setting, it will appear the code icon. When setting more than two parameters with code icon, LCD will just display one of the code icons. User should enter into the parameters setting interface to view the information or to restore the meter to the producer setting mode, and then reset the needed parameters.

4.5. Considerations:

4.5.1. Calibration times of meter rely on the sample, electrode performance and required accuracy. For high accurate measurement (≤±0.02pH), which should be calibrated immediately with high accurate standard buffer solution, for general accuracy measuring (≥±0.1pH), which can be used almost one week or long time once be calibrated.
4.5.2. The meter must be recalibrated in the following situations:

(a) New changed or unused electrode for a long time;
(b) After measuring acid (pH<2) or alkaline (pH>12) solution;

(c) After measuring solution which contains fluoride and concentrated organic solution;

(d) The solution’s temperature is much different with calibration temperature.

4.5.3. If do not want to use the temperature electrode, can press < ▲ > or < ▼ > key to adjust temperature to do manual temperature compensation.

4.5.4. The soaking solution in the protecting bottle of front pH electrode is to keep the glass bulb and junction activating. Loose the capsule, pull out the electrode and wash it in purified water before measuring. Insert the electrode and screw tight the capsule after measuring to prevent the solution leaking. If the soaking solution is turbid or moldy, please wash and change a new one at once.

4.5.5. The preparation of the soaking solution: take 25g analytical pure KCL, dissolved with purified water and dilute to 100mL. Electrode should avoid soaking in purified water protein solution and acid fluoride solution for a long time as well avoid getting touching with organic silicon lipidic matters.

4.5.6. To calibrate the meter with the given value pH buffer solution, the pH value of the standard buffer solution must be reliable so as to improve the accuracy. Buffer solution should be changed in time after many times using.

4.5.7. Always keep the meter clean and dry, especially for the socket of meter and electrode, otherwise it may lead to an inaccurate measurement or invalidity. To clean and dry them with medical cotton with dehydrated alcohol if there are any dirty.

4.5.8. The sensitive glass bulb in the front of combination electrode should not touch with hard things, any broken and rough will make the electrode invalidity. Before and after measuring, the electrode should be washed with purified water, and dry electrode after washing, don’t clean glass bulb with tissue for it will effect stability of electrode potential and enlarge response time. The electrode should be washed many times for removing the sample stuck on the electrode, or wash with suitable solvent then clean the solvent with purified water after measuring sticky sample.

4.5.9. An electrode be used for a long time, or measured solution which contains a polluting solute easily for the sensitive bulb, or a substance resulting in jam at the junction, the electrode will be getting passivated, its sensitivity will decrease and its response is getting slow, the reading are not correct. It could adopt the following method for various cases:
(a) The glass bulb is contaminated and aging: Put the electrode into 0.1mol/L dilute hydrochloric acid (Preparation: diluted 9mL hydrochloric acid to 1000mL with purified water) for 24h. Rinse it with purified water, then dipped it into the electrode dipping solution for 24h. If the passivation is serious, then user can also put the bulb of electrode into 4% HF (hydrofluoric acid) solution for 3 to 5 seconds, rinsing it with purified water, and dipped it in the electrode soaking solution for 24h to renew it. Note: the electrode activation solution is available in our company.

(b) Wash for contaminated glass bulb and junction:（FYI）

Contamination                           Abluent
Inorganic metal oxide                diluted acid less than 1mol/L
Organic lipidic matter                dilute washing（weak alkaline）

Resin macromolecule matter              dilute alcohol, acetone, ether

Proteinic haematocyte sediment      Acidic enzymatic solution（such as dried yeast）

Kinds of paint                      dilute bleacher, peroxide

4.5.10. pH electrode using period is about 1 year, but its life will be shortened if using condition is poor or incorrect maintenance. So it should be replaced immediately after electrode become aging or invalid.

4.5.11. When it appears an abnormal reading when calibration or displaying, please set P8 as “ON” to restore the meter to producer setting mode, and then to calibrate and measure again.

4.5.12. Do not unplug the adapter when the meter is power on, only to unplug when the meter is power off.

4.6. The Self-diagnose Information:

When using, there might appear the following icons. This is the meter’s self-diagnose information, which can help to know some information about the meter or the electrode when using:

[image: image71.png]Picture (4-6)



[image: image72.png]


4.6.1. The stable icon “          ” pH or the“         ” pH — this is the icon appeared when the value has surpassed the measuring range. There will also appear such signs when the electrode is not well connected with the meter or when the electrode is not immerged into the solution. This is a normal phenomenon.

[image: image73.png]


4.6.2. “        ” — Electrode zero potential to be exceeded  (<-60mV or >60mV)

[image: image74.png]


4.6.3 “        ” — Electrode slope to be exceeded（< 85% or >105%）

[image: image75.png]


[image: image76.png]


When appear “        ” or “        ”, the meter can not work, please take the following check:

(a) Check if the electrode bulb has air bubble, if has, please shake it hardly.

(b) Check the quality of buffer solution, if it goes bad or the value has biggish error.

(c) Set the meter to producer setting mode（for details see P14 item 4.4.10.）, then recalibrate it.

If still can not recover the normal state after doing the above checking, please replace a new pH electrode.

5. mV and ORP Measurement:

[image: image77.png]



[image: image78.png]


5.1. Sample Test:

[image: image79.png]OIOIO)



[image: image80.png]


[image: image81.png]


[image: image82.png]


5.1.1. Press <     > to electrify meter, press < MODE > to switch to     , then press < UNIT > key to choose      →      →      .

[image: image83.png]TR

—



(a)       — This is the measuring mode of electrode potential. mV value   and the temperature will appear at the same time.

[image: image84.png]{

"H

~—



(b)       — This is the ORP electrode measuring mode. ORP measurement has no temperature compensation, so there is no temperature appeared in this mode.
[image: image85.png]


(c)       — This is a special ORP measuring mode. It can show the ORP potential against the standard hydrogen electrode, EH value is related to reference electrode. The reference electrode potential in the producer setting is 204.6mV（3.5mol/L KCl, standard potential of Ag/AgCl reference electrode）. It is assorted with the ORP electrode supplied by our company. If choose other ORP electrode, please check the standard potential value according to its specifications and set the parameters in the meter’s parameter setting P1 (see P19 item5.3.2.);

Note: — ORP is the abbreviation of “ Oxidation-Reduction Potential ”, signify the Oxidation-Reduction potential of the solution, ORP is the index about the ability of the aqueous solution in the Oxidation-Reduction reaction, its unit is mV.

[image: image86.png]


5.1.2. Connect ORP electrode or ion electrode, immerge into the sample solution, slowly stir and then still. When there appear a “     ” and a stable reading, that is the ORP value or the potential of the ion electrode.

5.1.3. If it is a combination ORP or ion electrode, user just need to insert it into “ pH/mV/ISE ” socket, and if it is not combination type, then should also need to choose a proper reference electrode to connect to the “ REF ” socket and measure with both of the two electrodes.

5.2. Considerations:

5.2.1. The meter does not need to be calibrated when do ORP measurement. But if it is interrogative for the ORP electrode quality or the measuring result, it’s better to measure the mV value with the ORP standard solution to identify the accuracy of the ORP electrode or the meter.

Our company can supply the following ORP standard solution: 222mV(25℃)±15mV ORP standard solution (warranty 1 year), the inner solution of reference electrode is 3.5mol/L KCl.

Chart（5-1）

	℃     mV
	℃     mV
	℃     mV
	℃     mV

	10     242
	20     227
	30     215
	38     205

	15     235
	25     222
	35     209
	40     201


5.2.2. The clean and activation of ORP electrode: the surface of the platinum will be polluted after the ORP electrode has been using for a long time, which will lead to inaccurate measurement and slow response. It could adopt the following method for activation:

(a) For mineral pollution, immerge the electrode into 0.1mol/L dilute hydrochloric acid for 30 minutes, then wash with purified water, and then immerge into the electrode soaking solution for 6 hours before using.

(b) For the organic oil stain pollution, wash the platinum surface with detergent, and then immerge it into the electrode soaking solution for 6 hours before using.

(c) For the seriously polluted platinum with oxidized membrane, polish the surface with toothpaste, then wash it in purified water, and then immerge it into the electrode soaking solution for 6 hours before using.

5.3. Parameter Setting:
5.3.1. mV and ORP measurement parameter setting schedule (chart (5-2))

Chart（5-2）

	Prompt Mark
	Parameter Setting Items
	Code 
	Parameters

	P1
	Reference electrode potential setting
	[image: image87.png]



	0 to 999.9mV

	P2
	Time for timing measuring setting
	[image: image88.png]



	0 to 99 minutes

	P3
	Restore to producer setting
	
	OFF-On (shut-set)


5.3.2. Reference electrode potential setting:

[image: image89.png]


[image: image90.png]


[image: image91.png]


[image: image92.png]


(a) Depress < MODE > key in the mode     ,     ,

meter enters into mode P1, see picture（5-1）.

(b) Press < UNIT > key, the number will move rightward in turn and twinkle, press < ▲ > or < ▼ > to alter the number when the number twinkling.
(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to confirm and return to measuring mode.

(d) The producer setting is 204.6mV.

[image: image93.png]


[image: image94.png]


Note: If a potential value different from the producer setting value has been set, there will appear “      ”icon in the upper right corner of the LCD in      measuring mode.

5.3.3. Timing measuring setting (P2): 

Please see P13 item 4.4.6.

5.3.4. Restore to producer setting（P3）:

Please see P14 item 4.4.10.

6. Ion Concentration Measurement:

[image: image95.png]U-EF




[image: image96.png]


6.1. Preparation Work:

[image: image97.png]


6.1.1. Switch in the power, press <     > to turn on the meter; short press < MODE > key switch to     measuring mode; 

6.1.2. Install the electrode holder and adjust it, electrify the model 901 stirrer for using;

6.1.3. Choose the proper ion selective electrode, reference electrode and temperature electrode, connect them to the meter;
6.1.4. Choose measuring unit: press < UNIT > key to choose pX, mol/L, mg/L or ppm;

Note:
(a) The symbol of pX unit vary from different ion selective electrodes, such as pNa, pCa and so on, so there is no displayed symbol when choosing pX unit on the surface of LCD.

[image: image98.png]


(b) When choosing the unit pX, ppm or mg/L, LCD display mode will show as picture（6-1）, the measuring reading is 168.6ppm displaying in the main display part;

[image: image99.png]


[image: image100.png]oC
)

Picture (4-8)




(c) When choosing unit mol/L, LCD display mode will show as picture（6-2）, the measurement value adopts the combination form, the “       ” signifies 10-2 , and the complete measuring reading is 1.68×10-2 mol/L;

6.1.5. Choose variety of ions: there are 10 kinds of common ions which have been set in meter, in addition, molecular weight and ionic valence also can be set in to self-define ions. For detailed operation please see P21 item 6.2.2. ion variety selection P1 and P21 item 6.2.3. self-define ion setting P2;

6.1.6. Prepare two kinds of calibration solution, and setup to the meter separately. The detailed operations please see P21 item 6.2.4. calibration solution setting P3;

6.2. Parameter Setting:
6.2.1. Ion concentration parameter setting schedule (chart (6-1))

Chart（6-1）

	Prompt Mark
	Parameter Setting Items
	Code
	Parameters

	P1
	Ion variety selection
	
	F-, Na+, Cl-, NH4+, NO3-, Ca2+, Br-, I-, K+, Cu2+

	P2
	Self-define ion setting
	[image: image101.png]I/

P UUGE

=39 O





	Self-define

	P3
	Calibrate solution setting
	
	C1 and C2

	P4
	Timing measuring setting
	[image: image102.png]ISE





	0 to 99 minutes

	P5
	Restore to producer setting
	
	OFF-On（shut-set）


6.2.2. Ion variety selection（P1）

[image: image103.png]mV




[image: image104.png]mV




[image: image105.png]ORP




(a) Depress < MODE > key to enter into P1 mode, show as picture （6-3）, there are 10 kinds of common ions can be selected. “ F- ” is flashing, means it has been selected, and the upper right corner “     ” means the valence of “ F- ”, and the lower right corner “      ” means molecular weight of F-, press < ▲ > or < ▼ > key to move the flashing ion symbol to choose other ions;

(b) Press < MODE > key to enter into the next parameter setting or press < ENTER > key to confirm and return to measuring mode.

[image: image106.png]


6.2.3. Self-define ion setting（P2）
[image: image107.png]


(a) Short press < MODE > key in mode P1 to enter into mode P2, show as picture（6-4）, “      ” means the “Self-define” interface has been shut off.

[image: image108.png]ISE




[image: image109.png]


[image: image110.png]mV




(b) Press <UNIT> key to open the interface of “Self-define”, the interface show as picture（6-5）, the upper right corner “   ” means the ion valence setting, and the main shown area “           ” means the molecular weight setting. Now take the Ag+ (molecular weight 107.9) as an example to set up; 

[image: image111.png]


[image: image112.jpg]


[image: image113.png]


(c) Press < UNIT > key to shift, the order is “   ” is flashing in the upper right corner→ “   ” in the main display moving rightward in turn and flashing, press <▲> or < ▼ > to change the number when the number is flashing, revise the number in the main displaying are to be “        ”;

(d) Press < MODE > key to enter into the next parameter setting or press < ENTER > to confirm and return to measuring mode.
Note: after setting the self-define mode, the LCD will not display the ion symbol, indicates this is the interface under the self-define mode; at this time if set any kind of ion in the mode of P1 (such as “Na+”), the  self-define interface will turn off automatically, and the measuring interface will display “Na+” symbol;

6.2.4. Calibration solution setting（P3）

[image: image114.png]Picture (4.7)




6.2.4.1. The calibration solution setting for the unit pX, ppm or mg/L. Take the standard solution C1=0.019ppm, C2=1.900ppm of F- as an example:

(a) Short press < MODE > in mode P2 to enter into P3 mode; LCD show as picture（6-6）, C1 means the first solution setting.
[image: image115.png]Picture (4-5)



[image: image116.png]717,
Ll

Picture (4-1)




[image: image117.jpg]


[image: image118.png]7 S

| pH [mV ISE] 101010 Date
[ORP| EH | LIC:000 Time
at—ef 5] I=] I ] :,";"r:gﬁt
L L.l. o
o OB iGgaE
@@@ @ & EEHR




[image: image119.png]


(b) Press < UNIT > key to shift, the order is “   ”（twinkle）→“   ”（twinkle）→“   ”（twinkle）→“   ”（twinkle）→“·”（twinkle）; press < ▲ > or < ▼ > to change the number when the numbers are twinkling. When the decimal point is twinkling, press < ▲ > key to move rightward. User can set up the number as “          ” ppm according to the above operation;

[image: image120.png]


[image: image121.png]


(c) Press < ENTER > to confirm and enter into “ C2 ” setting. The interface show as picture (6-7), C2 refers to the second solution setting; to operate the same as the above item（b）to set up the number as “          ”, press < ENTER > key to confirm and return to measuring mode.

[image: image122.png]“““““

Picture (6-10)



6.2.4.2. Calibration solution setting when unit is mol/L, take the calibration solution C1=1.50×10-6mol/L and C2=1.50×10-4mol/L of the F- as an example:

(a) Under P2 mode to short press < MODE > key and enter into P3 mode, LCD interface show as picture（6-8）, display value is 1.00 ×10-5, C1 means the first calibration solution;

[image: image123.png]


[image: image124.png]


[image: image125.png]“““““

Picture (6-11)



[image: image126.png]


[image: image127.png]


[image: image128.png]


[image: image129.png]H



[image: image130.png]I H



[image: image131.png]


[image: image132.png]Picture (6-7)



(b) Press < UNIT > key to shift, the order is “   ”（flashing）→“   ”（flashing）→“   ”（flashing）→“       ”（flashing）;When the number or “       ” flashing, press < ▲> or < ▼> key to change the number. According to the above operation, set up the main display area as “      ”, and set the index to be “       ”;
[image: image133.png]


(c) Press < CAL > key to confirm and enter into “     ” setting, LCD interface show as picture（6-9）, C2 means the second solution setting; According to the above same [image: image134.png]-



item（b）operation, set the main display area as “       ”, set the index to be “       ”, then press < ENTER > key to confirm and return to the measuring mode.

[image: image135.png]-



6.2.5. Time for timing measuring setting（P4）

Under P3 mode, short press < MODE > key to enter into the P4 mode. For the other operations please refer to the P13 item 4.4.6.

6.2.6. Restore to producer setting（P5）

Under P4 mode, short press < MODE > key to enter into the P5 mode, for other operations please refer to the P14 item 4.4.10.

6.3. Electrode Calibration:

6.3.1. Press < UNIT > key to select the unit, such as choosing the ppm unit;
6.3.2. Selecting the ion variety in P1 parameter setting, such as choosing F-;

6.3.3. Set two calibration solution in the parameter setting P3, such as C1=0.019ppm C2=1.900ppm;

[image: image136.png]


[image: image137.png]q”:
22



6.3.4. Press < CAL > key and the meter enters into calibration mode, 0.019ppm displaying on the LCD and flashing “     ” indicate entering into the first point calibration. The calibration solution is 0.019ppm, see picture（6-10）;

[image: image138.png]Picture (6-2)



[image: image139.png]


[image: image140.png]-



6.3.5. Wash the electrodes（ISE + reference electrode + temperature electrode）with purified water, and insert into the C1 standard solution, after shaking the electrode holder short moment, then still it and wait for 1 to 2 minutes. Again press < CAL > key when “     ” appears, LCD will display flashing 0.019ppm, the calibration is finished after several seconds, and LCD display 1.900ppm and flashing “    ”, indicating enters into the second point calibration, the standard solution is 1.900ppm, see picture（6-11）;
[image: image141.png]Picture (§-5)



[image: image142.png]neE



[image: image143.png]neE



6.3.6. Wash the electrodes with purified water, and insert it into the standard solution C2, shaking the electrode holder short moment, then still and wait for 1 to 2 minutes. After the “     ” icon appears, again press < CAL > key, LCD displays the flashing 1.900ppm, then the calibration is finished after several seconds and return to measuring mode, to display stable measurement value and      ,      two calibration indication icons.

Note: It only display the standard solution which has been set by user and do not display measuring reading while meter’s calibrating. But press the < UNIT > key at that time, can enter into mV mode, showing actual measuring mV value, and also can finish the calibration procedure under the mV mode.

6.4. Sample Test:

[image: image144.png]Picture (6-4)



6.4.1. Wash the electrode with purified water and dry, insert it into the solution and shaking the electrode holder for short moment, then still and take reading when the measurement is stable and “     ” icon appears, then it is the ion concentration of solution;

6.4.2. Press < UNIT > key, it can regularly display the ion concentration of other unit.

6.5. Considerations:

6.5.1. Because there are many varieties of ions, hence this meter does not equip ISE, and users need to choose ion electrode according to their demand. When selecting an ion electrode, please specially consider the detection limit of the ISE in order to accord with the using requirement of customers.
6.5.2. The meter equips with 1pc of model 6213 double salt bridge reference electrode, which the inner and outer reference solution can be replaceable. The user may accord to the different character of ion electrode to select the right inner and outer reference solution. Commonly, the inner reference solution adopted KCl solution, and the outer reference solution is adopted the other electrolyte solution which no disturbing on ion measurement.

6.5.3. According to the characteristics of ISE to choose the correct total ionic strength regulator will make measurement to be more stable, particularly the concentration of measured ion is near to detection limit of the ISE. When using, should pay attention to add the equal volume of ionic strength regulator to the calibration solution and measuring solution which have the same volume.
6.5.4. When making a test with ISE, it will be more stable in the stirring state, particularly the concentration of measured ion is near to detection limit of the ISE, but the stirring speed should not be too fast.

6.5.5. User need to prepare two sorts of suitable calibration solution and the concentration of it should be in the measuring range of ion electrode, avoid to choose the calibrate solution under the detection limit of ion electrode. In order to ensure the accuracy of measuring, the measuring value of sample solution should be between the concentration of this two calibration solution.
6.5.6. The concentration of C1 and C2 two kinds of calibration solution should be from dilute to concentrate. During the calibrating and measuring, if the solution is from dilute to concentrate, using of electrode needn’t wash, just throw off the water on it. If it is from concentrate to dilute, the electrode should be washed repeatedly, otherwise it will affect its accuracy.

6.5.7. Calibration solution of ion electrode does not has buffer action and easy to be polluted particularly to the low concentration of solution. Hereby, carefully prepare calibration solution to be necessary and pay attention to the clearance on using to avoid the stain.

6.5.8. Since the meter is without identification on calibration solution of ion electrode, so, it’s necessary to accord with “C1” or “C2” icon which flashing on the LCD to insert the electrode into the matched calibration solution. Otherwise, it will have a lager error, even can’t use.

6.5.9. Meter adopts two-point calibration method, the meter displayed the data is not accurate before finishing the calibration.
6.5.10. When it appears an abnormal reading when calibration or displaying, please set P5 as “ON” to restore the meter to producer setting mode, and then to calibrate and measure again.

6.5.11. Do not unplug the adapter when the meter is power on, only to unplug when the meter is power off.

7. Using Guide on Model 901 Intelligent Stirrer: 
[image: image145.png]Picture (6-3)





7.1. Technical Specifications:

	Timing range
	0 to 2300 rpm（empty load）

	Max. working current
	150mA

	Max. power consume
	0.9W

	Working area diameter
	Ф100mm

	Maxi. stirring capacity
	1000mL

	Power
	DC6V power adapter or 4 pcs of AA batteries

	Size and weight
	100 × 146 × 48 mm; weight 340g


7.2. Functional Features:

7.2.1. Adopts microprocessor chip design, can memory the rotation speed;

7.2.2. ABS plastic housing, transparent coping, beautiful appearance, and slim compact;
7.2.3. Reliable waterproof design（waterproof rating IP57）;

7.2.4. Alternative choice from AC/DC, can be used in field, on-site and incubator;
7.2.5. Warning function for low battery power（<5V）.
7.3. Using Method:

[image: image146.png]


7.3.1. Plug the DC6V power adapter, or open the bottom cover to load 4 unit of AA batteries;
7.3.2. Press <     >key（bright green indicator light）, means the stirrer is electricity connected; press < ▲ > and < ▼ > key to adjust the rotation speed; press < RPM > key (bright green indicator light), stirrer is working accordance with original stored rotation speed; again press < RPM > key to turn off stored rotation speed and renew to the adjusted well rotation speed. Hereby, it can realize alternatively two kinds of rotation speed by turning on and turning off < RPM > key.
7.3.3. Storing rotation speed: depress < RPM > key for 3 seconds with green indicator is flashing, showing this rotation speed is stored. It will appear after turning on the meter by pressing the < RPM > key next time.

7.3.4. The stirrer is divided 100 grades from zero to the highest rotation speed, press < ▲ > key for 100 times, can achieve to the highest rotation speed and can do accurate adjustment when using. However, if depress < ▲ > or < ▼ > key, can quickly increase or decrease the rotation speed.
7.4. Considerations:

7.4.1. If the bottom plane of beaker is uneven, the vibration possibly will produced even can’t stir, encountering this situation, should replace a more qualified beaker.

7.4.2. When use of battery, if electricity is low than 5V, the red indication light which on the bottom of the faceplate will be bright, showing the original set of standard rotation speed will reduce, and should replace new batteries at once. However, this stirrer will still work at that time, only speed is lower, press < ▲ > key can increase speed, if user do not care about the accuracy of original set standard speed, even red lights bright stirrer still able to work, until the battery depletion of electricity.
7.4.3. In the situation of zero rotation speed, do not depress < RPM > key. Otherwise, it’s possible to set up as zero rotation by mistake. Press < RPM > key and the bright green do not light, and rotation speed is zero, please just press < ▲ > key to renew the setting.
8. Meter’s Complete Kit:
[image: image147.png]Picture (4-2)




[image: image148.png]


8.1. Model MP523 pH/ISE meter

8.2. Model 602 flexible electrode holder
8.3. Model 901 intelligent stirrer
8.4. 201T-M plastic pH/ATC three-in-one combination electrode

8.5. LabSen2503 glass pH combination electrode
8.6. MP500 temperature electrode
8.7. Model 6213 double salt bridge reference electrode

8.8. pH standard buffer solution（pH4.00, pH7.00 and pH10.01/50mL）
8.9. 9V power adapter（meter use）

8.10. 6V power adapter（stirrer use）
8.11. B628 stirring bead

8.12. RS-232 communication cable

8.13. MP523 communication software disk

8.14. Operation manual
8.15. Brief operation instruction

9. Warranty:

[image: image149.png]



9.1. We warrant this meter to be free of charge maintain, replace the parts or products under normal using circumstances, from purchased time within one year caused by manufacturing bad and unable to work.
9.2. Attached other electrodes are not belong to this warrant range, besides the temperature electrode. But, if the newly purchased electrode went wrong without using, it’s free of charge to maintain and replace.
9.3. The above warranty is not apply to defects resulting from action of user such as misuse, improper wiring, operation outside of specification, improper maintenance or repair, or unauthorized modification.
[image: image150.png]



Chart І Meter Parameter Setting Schedule
	Mode
	Prompt Mark
	Parameter Setting Items
	Code
	Parameter

	pH
	P1
	pH buffer solution series selection
	[image: image151.png]ISE (2100
\_ 4
)

(- e
(Y

< )
C 0 ATC

© ®

Picture (6-1)





	USA (Europe & U.S.A. series)
NIS (NIST series)
CH (China series)
CUS (Self-define)

	
	P2
	Purified water pH temperature compensation setting
	[image: image152.png]Erd




	OFF-On (shut-set)

	
	P3
	Purified water added with ammonia pH temperature compensation setting
	[image: image153.png]En





	OFF-On (shut-set)

	
	P4
	Time for timing measuring setting
	
	0 to 99 minutes

	
	P5
	Temperature unit setting
	
	℃ ℉

	
	P6
	Date setting
	Date
	Month / Day / Year

	
	P7
	Time setting
	Time
	Hour / Minute

	
	P8
	Restore to producer setting
	
	OFF-On (shut-set)

	mV
	P1
	Reference electrode potential setting
	[image: image154.png]En





	0 to 999.9mV

	
	P2
	Time for timing measuring setting
	[image: image155.png]ORP





	0 to 99 minutes

	
	P3
	Restore to producer setting
	
	OFF-On (shut-set)

	Ion
	P1
	Ion variety selection
	
	F-, Na+, Cl-, NH4+, NO3-, Ca2+, Br-, I-, K+,Cu2+

	
	P2
	Self-define ion setting
	[image: image156.png]



	Self-define

	
	P3
	Calibrate solution setting
	
	C1 and C2

	
	P4
	Time for timing measuring setting
	[image: image157.png]



	0 to 99 minutes

	
	P5
	Restore to producer setting
	
	OFF - On (shut-set)


Chart Ⅱ Meter Restore to Producer Setting Schedule 
	Mode
	Prompt Mark
	Parameter Setting Items
	Producer Setting Content
	Non Producer Setting Content Icon

	pH
	P2
	Purified water pH temperature compensation setting
	OFF
	[image: image158.png]U-EF





	
	P3
	Purified water added with ammonia pH temperature compensation setting
	OFF
	[image: image159.png]




	
	P4
	Time for timing measuring setting
	0 second
	[image: image160.png]




	mV
	P1
	Reference electrode potential setting
	204.6mV
	[image: image161.png]




	
	P2
	Time for timing measuring setting
	0 second
	[image: image162.png]




	Ion


	P2
	Self-define ion setting
	OFF
	[image: image163.png]




	
	P6
	Time for timing measuring setting
	0 second
	[image: image164.png]H






Chart Ⅲ Code Icon and Abbreviation Schedule
[image: image165.png]



 Shanghai San-Xin Instrumentation, Inc.
Add: 4/F. Building No.16, No. 481 Guiping Road, Shanghai, China. 200233
Tel: +86-21-63362480               Fax: +86-21-64956880

Website: www.shsan-xin.com        E-mail: wxmab@shsan-xin.com
� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���



































�





Setting 1.090 pH according to item (b)





�





�





�











� EMBED PBrush  ���








� EMBED PBrush  ���








�





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���











3


4


5


5


6


7


7


8


8


8


9


10


10


14


16


17


17


18


19


19


19


20


23


24


24


25


25


26


26


26


27


27




















� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





�





� EMBED PBrush  ���








� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





�





�





�





� EMBED PBrush  ���





�





ON


OFF





� EMBED PBrush  ���





� EMBED PBrush  ���





�





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





�





�





�








� EMBED PBrush  ���








�





� EMBED PBrush  ���








� EMBED PBrush  ���





� EMBED PBrush  ���� EMBED PBrush  ���





� EMBED PBrush  ���





�





�





�� EMBED PBrush  ���





�





�





�





�





Press < ENTER > two times, return to measuring mode (one point calibration mode)





Press < ENTER > two times, return to measuring mode (two point calibration mode)


Picture（4-3）





Press < ENTER > two times, return to measuring mode (three point calibration mode)





�





�





�





Press <CAL> key








Press <CAL> key








�





�





�





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���








�





� EMBED PBrush  ���





� EMBED PBrush  ���








� EMBED PBrush  ���





� EMBED PBrush  ���





Setting 3.776 pH according to item (b)





�





�





�





�





� EMBED PBrush  ���





�





�





ON


OFF  





� EMBED PBrush  ���





� EMBED PBrush  ���





�


























� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���











1unit


1pc


1unit


1pc


1pc


1pc


1pc


1btl/each


1pc


1pc


3pcs


1pc


1pc


1pc


1pc





� EMBED PBrush  ���





ON


OFF























�















































�





























�














� EMBED PBrush  ���





� EMBED PBrush  ���





�





� EMBED PBrush  ���




















�





�





�





� EMBED PBrush  ���








� EMBED PBrush  ���





ON


OFF








�





�











�





�





�











�





�
































� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





�








� EMBED PBrush  ���





� EMBED PBrush  ���








�





� EMBED PBrush  ���





�





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





�





�





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





�





� EMBED PBrush  ���





� EMBED PBrush  ���





�





��





�





�





� EMBED PBrush  ���





�





Setting 7.413 pH according to item (b)





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���





�





ON


OFF











Calibration











� EMBED PBrush  ���











End












































�





�





�





� EMBED PBrush  ���





� EMBED PBrush  ���





� EMBED PBrush  ���











Volt Reference





Reference electrode potential setting





� EMBED PBrush  ���











Pure-2





Purified water added with ammonia pH temperature compensation setting





� EMBED PBrush  ���











Pure-1





Purified water pH temperature compensation setting





� EMBED PBrush  ���











Custom





Self-define solution





� EMBED PBrush  ���











NIST





NIST series standard





� EMBED PBrush  ���











USA





Europe & U.S.A series standard





� EMBED PBrush  ���











China





China series standard





� EMBED PBrush  ���











Solution





Standard solution









































� EMBED PBrush  ���











Oxidation-Reduction Potential



































Code and Abbreviation





English





Explanation





� EMBED PBrush  ���











Ion Selectivive Electrode















31
MP523 Version2.0 2018-6-5

[image: image166.png]ORP




[image: image167.png]ISE




_1273666235

_1285236669

_1285238294

_1285477045

_1285236828

_1285238211

_1279694343

_1285234949

_1285235013

_1285235195

_1279694361

_1273666346

_1274266675

_1275116577

_1273666664

_1273666284

_1273574680

